## Slide 1
Supplemental File

## Slide 2
60s ave
Measurement of the inhalation
O2_raw [%]
(Time series data)
Switch the ventilation circuit
tE [sec]
(Gas collection)
tI [sec]
(Rebreathing)
CO2_raw [%]
(Time series data)
T1
60s ave
60s ave
Measurement:
- Gas temperature [degC]
- Gas humidity [%RH]
- Atmospheric pressure [hPa]
Airway pressure [mmHg]
*differential pressure
(Time series data)
O2 sensor
CO2 sensor
Airway pressure
60s ave
60s ave
Time
T2
T3
T4
FiO2_est [%]
(Time series data)
Measurement of the inhalation
Measurement of the exhalation
Switch the ventilation circuit
T2 ~ T3 ave
Exhalation airway pressure
Inhalation airway pressure
Fig. 1

## Slide 3
Figure 1. Diagram of Experimental Work Flow
The time series data of concentrations of oxygen and carbon dioxide as well as intra-circuit airway pressure were recorded continuously.
T1: the ventilator ports were switched from the inhalation to exhalation sides and the sensors were placed into the gas flow of exhalation.
T2: starting collection of the exhalation gas into the soft bag and replacing the gas bag to an inlet of the cylinder pump (R measurement).
T3: the ventilator ports were switched from the exhalation to inhalation sides and the sensors were placed into the gas flow of inhalation.
T4: the end of time series data.
Taking into account of the drift of O2 sensor, FiO2 estimation curve was drawn linearly from T1 to T2.

## Slide 4
Atmospheric pressure [hPa]
Airway pressure [mmHg]
*differential pressure
(Time series data)
O2_raw [%]
(Time series data)
CO2_raw [%]
(Time series data)
FuncTimeSeriesDataCalc
Pressure rise rate Ex
Gas temperature [degC]
Gas humidity [%RH]
tE [sec]
tI [sec]
FuncViVeRatioCalc
R_humid
R_dry
FeO2_corrP_ave [%]
- ATPS
FeCO2_corrPcorrO2_ave [%]
- ATPS
FiO2_est_corrP_ave [%]
- ATPD
RQ calculation with the measured R_humid
RQ = FeCO2_corrPcorrO2_ave / (R_humid * FiO2_est_corrP_ave - FeO2_corrP_ave)
RQ
Calculate vol% of humidity
(1) Calculate “Saturated vapor pressure [hPa]” from “Gas temperature [degC]”
(2) Vapor pressure [hPa] = Saturated vapor pressure * Gas humidity / 100
(3) Vol% of humidity in the gas collection bag: FcVapor = ( Vapor pressure / Atmospheric pressure ) * 100
FeO2_raw_ave [%]
- Exhalation airway pressure
- Humid
FeCO2_raw_ave [%]
- Exhalation airway pressure
- Humid
- Affected by O2 level
FcVapor [%]
Correction considering pressure
FeO2_corrP_ave [%] = FeO2_raw_ave / Pressure rise rate EX
Correction considering pressure
FeCO2_corrP_ave [%] = FeCO2_raw_ave / Pressure rise rate Ex
FeCO2_corrP_ave [%]
- ATPS
- Affected by O2 level
Correction considering the interaction between CO2 and O2
Correction coefficient: a = -0.00072
FeCO2_corrPcorrO2_ave [%] = FeCO2_corrP_ave / (1 + a * FeO2_corrP_ave)
Fig. 2

## Slide 5
Figure 2. Correction Diagram and Calculation Algorithm
The figure shows a scheme of the correction diagram and the calculation algorithm of respiratory quotient. Intra-circuit pressure, O2 concentration, and CO2 concentration in the ventilator circuit and atmospheric pressure were monitored. Gas temperature and humidity were measured at the exhaust port of the ventilator and the exhalation gas was collected into the bag through the port. The calculation algorithm of molecular ratio of inhalation to exhalation, named “R” in our main text, is seen on the left half of the scheme. FiO2, FeO2, and FeCO2 were first corrected by the intra-circuit pressure and set at the atmospheric pressure level. Since CO2 sensor is affected by the concentration of O2, the number of FeCO2 was corrected with the number of FeO2.

## Slide 6
FuncTimeSeriesDataCalc
Atmospheric pressure [hPa]
O2_raw [%]
(Time series data)
CO2_raw [%]
(Time series data)
Average (60 sec, ~T1)
- for the inhalation (initial)
Average (60 sec, ~T4)
- for the inhalation (post procedures)
Average (60 sec, ~T2)
- for the exhalation
FiO2_raw_ave_initial [%]
- Inhalation airway pressure
- Dry
FiO2_raw_ave_post [%]
- Inhalation airway pressure
- Dry
FeO2_raw_ave [%]
- Exhalation airway pressure
- Humid
FiO2 estimation considering drift of the O2 sensor output
(1) Slope: a = (FiO2_raw_ave_post - FiO2_raw_ave_initial) / (T4 - T1)
(2) FiO2_est = a * (elapsed time - T1) + FiO2_raw_ave_initial
FiO2_est [%]
(Time series data)
- Inhalation airway pressure
- Dry
Average (60 sec, ~T2)
- for the inhalation
FiO2_est_ave [%]
- Inhalation airway pressure
- Dry
Correction considering pressure
(1) Total pressure In [hPa] = Atmospheric pressure + ( Inhalation airway pressure / 760 ) * 1013.25
(2) Pressure rise rate In = Total pressure In / Atmospheric pressure
(3) FiO2_est_corrP_ave [%] = FiO2_est_ave / Pressure rise rate In
FiO2_est_corrP_ave [%]
- ATPD
Airway pressure [mmHg]
*differential pressure
(Time series data)
Average (60 sec, ~T1)
Average (T2 ~ T3)
Inhalation airway pressure [mmHg]
*differential pressure
Exhalation airway pressure [mmHg]
*differential pressure
Calculate total pressure of exhalation
Total pressure Ex [hPa] = Atmospheric pressure + ( Exhalation airway pressure / 760 ) * 1013.25
Total pressure Ex [hPa]
Calculate rise rate of the pressure of exhalation
Pressure rise rate Ex = Total pressure Ex / Atmospheric pressure
Pressure rise rate Ex
Average (60 sec, ~T2)
- for the exhalation
FeCO2_raw_ave [%]
- Exhalation airway pressure
- Humid
- Affected by O2 level
Fig. 3

## Slide 7
Figure 3. Data Collection Algorithm
The time series data of concentrations of oxygen and carbon dioxide as well as intra-circuit airway pressure were recorded continuously. T1 to T4 in this figure are referred from Figure 1 in this supplemental file.

## Slide 8
tE [sec]
- Time representing inhalation volume
- During the gas collection into the bag
- Dry
R_humid
FcVapor [%]
FuncViVeRatioCalc
tI [sec]
- Time representing exhalation volume
- During the rebreathing from the bag
- Humid
Correction considering permeation of humidity through the gas collection bag
Correction coefficient: a = -0.0857
tI_adjusted = tI / (1 + a * FcVapor/100)
tI_adjusted [sec]
- Time representing exhalation volume
- During the rebreathing from the bag
- Humid
Humid R calculation
R_humid = VI_dry / VE_humid = tE / tI_adjusted
Correction considering humidity
tI_adjusted_dry = tI_adjusted * (1 - FcVapor/100)
tI_adjusted_dry [sec]
- Time representing exhalation volume
- During the rebreathing from the bag
- Dry
Dry R calculation
R_dry = VI_dry / VE_dry = tE / tI_adjusted_dry
R_dry
Fig. 4

## Slide 9
Figure 4. Collection Algorithm of R
The molecular ratio of inhalation to exhalation is referred as R in our main text. R was measured by dry inhalation and humid exhalation in our animal model (R-humid). Gas temperature and humidity were measured at the exhaust port of the ventilator and the exhalation gas was collected into the bag through the port. R dry was calculated from the numbers of humidity, temperature, and R-humid.