## Slide 1: Single-Molecule Localization Microscopy: Theoretical Basis and Practical Guide
“Widening the Lens” Microscopy Education program Vanderbilt University, Nashville, TN, USA Ulrike Boehm, Ph.D. ulrike.boehm@zeiss.com - @ulrike_boehm Corporate Research & Technology (CRT), Carl Zeiss AG

## Slide 2: Diffraction Limited
vs. Diffraction Unlimited
Microtubules, scale bar: 1 um

## Slide 3: The Diffraction Limit
Abbe et al. Arch. Mikroskop. (1873)
Abbe Equation
Lateral resolution
Axial resolution
Point Spread Function (PSF)
Notes: PSF – point spread function

## Slide 4: Circumventing the diffraction limit with…
Schermelleh et al. Nature Cell Biology (2019)
SIM
STED
SMLM
This lecture’s focus

## Slide 5: Major discoveries enabled by SMLM
Lelek et al. Nature Reviews Methods Primer (2021)

## Slide 6: Major discoveries enabled by SMLM
Lelek et al. Nature Reviews Methods Primer (2021)

## Slide 7: Overview
Basic principles 2D SMLM
Hardware for SMLM
Practical considerations: sample preparation, suitable dyes, linkage errors, and buffers
Processing, quantification, and interpretation of SMLM data
3D SMLM
Summary
Extra: New directions in SMLM
Extra: SRM as a multidimensional challenge

## Slide 8: Basic principles of 2D SMLM

## Slide 9: What is Single Molecule Localization Microscopy (SMLM)?
Widefield image
Betzig et al. Science (2006), Hess et al. Biophy. J. (2006), Rust et al. Nature Methods (2006)
Sequence of widefield images
Accumulated localizations
Superresolution image

## Slide 10: What is Single Molecule Localization Microscopy (SMLM)?
Image Acquisition
Image at t1
Image at t2
Image at t3
t1 < t2 < t3
Hardware
Sample Preparation
Processing
Betzig et al. Science (2006), Hess et al. Biophy. J. (2006), Rust et al. Nature Methods (2006)
Superresolution image
Widefield image

## Slide 11: What is Single Molecule Localization Microscopy (SMLM)?
In SMLM small subsets of individual emitters are randomly activated or switched ON/OFF in consecutive acquisitions.
If sparse enough to be identified as single molecule switching events, signals become spatiotemporally separated and are collected over several thousands of camera frames.
Raw data are computationally processed to detect single molecules and determine their center positions with nanometer precision dependent on the number of photons detected per individual emitter.
These are finally assembled through superimposition into a single-plane image.
Betzig et al. Science (2006), Hess et al. Biophy. J. (2006), Rust et al. Nat Methods (2006)

## Slide 12: Overview
Basic principles 2D SMLM
Hardware for SMLM
Practical considerations: sample preparation, suitable dyes, and buffers
Processing, quantification, and interpretation of SMLM data
3D SMLM
Summary
Extra: New directions in SMLM
Extra: SRM as a multidimensional challenge

## Slide 13: Hardware for SMLM

## Slide 14: Hardware for SMLM
Wide-field
TIRF
Schermelleh et al. Nature Cell Biology (2019)
Java Tutorial: Evanescent Field Penetration Depth

## Slide 15: Hardware for SMLM
Wide-field or TIRF with high NA objectives (15k-20k)
High power lasers (~25k)
Sensitive EMCCD (~35k) or sCMOS (~16k) cameras
Additional optics for 3D information
System upgrade: ~$80,000

## Slide 16: Hardware for SMLM
Ma et al. Scientific Reports (2017)
Total cost: $4,100.00

## Slide 17: Overview
Basic principles 2D SMLM
Hardware for SMLM
Practical considerations: sample preparation, suitable dyes, linkage errors, and buffers
Processing, quantification, and interpretation of SMLM data
3D SMLM
Summary
Extra: New directions in SMLM
Extra: SRM as a multidimensional challenge

## Slide 18: Practical considerations: sample preparation, suitable dyes, linkage errors, and buffers

## Slide 19: Photoswitching – ON/OFF switching

## Slide 20: Photoswitching – ON/OFF switching
~ps
Excitation
ON
OFF
Fluorescence
~ps
Conventional dyes
Temporal separation not possible
OFF
ON
Typical SMLM dyes
“Dark” state allows temporal separation for SMLM
Dark state decay rates: ~µs
Electron lifetime: ~ns
Notes: The off state can be a dark state, a state with a different fluorescent emission etc. – it only needs to be separate from the original state
Dashed line – not necessary reverse

## Slide 21: A good SMLM dye has…
High absorption coefficient and quantum yield to produce bright events.
Bleaching, brightness, ON-time, and ON/OFF duty cycle are all strongly affected by laser power and buffer composition.
Nahidiazar et al. PLOS ONE (2016)
Long ON-time to allow the collection of many photons, but not too long such that the data acquisition is slowed down.
Complete dark OFF-state that is long enough to prevent too many fluorophores from being ON simultaneously.
Exceptional resistance to bleaching since tens of thousands of images are collected at high laser power.

## Slide 22: Photoswitchable Fluorescent Proteins
Jradi et al. ACS Chem. Biol. (2019)
Photoactivation by decarboxylation (PA-GFP, PA-mKate, PA-mCherry)
Photoswitching by beta-elimination (Dendra, mEos, mClavGR2, mIrisFP, mMaple)

## Slide 23: Photoswitchable Fluorescent Proteins
Jradi et al. ACS Chem. Biol. (2019)
Photoconversion by cis-trans isomerization (Dronpa)
Photoconversion by hydration/dehydration (Dreiklang)

## Slide 24: Photoswitchable Fluorescent Proteins
Jradi et al. ACS Chem. Biol. (2019)
Notes: As with all fluorescence microscopy, the photophysical properties of the fluorophore labels are crucial for SMLM. These are summarized in Tables 1 and 2 and include the laser excitation wavelength (λex), the fluorescence emission maximum (λem), the extinction coefficient at λmax (ε), and the fluorescence quantum yield (Φ). Properties specific to fluorophores activated with light include: the fluorescence contrast upon photoactivation (Fon/Foff), the turn-on (t1/2on) halftime, the photobleaching halftime (t1/2PB), and the photon yield of individual fluorophores before bleaching (N). The maturation halftime (t1/2M) is an important consideration for fluorescent proteins. For caged dyes, a critical factor is the photoactivation quantum yield (ΦPA).

## Slide 25
Jradi et al. ACS Chem. Biol. (2019)

## Slide 26: Photoswitchable Dyes
Jradi et al. ACS Chem. Biol. (2019)
Conventional fluorophores (Alexa Fluor 488, Cy5, Cy3)

## Slide 27: Photoswitchable Dyes
Jradi et al. ACS Chem. Biol. (2019)
Spontaneously blinking (HM-SiR)
Photochromic rhodamines

## Slide 28: Photoswitchable Dyes
Jradi et al. ACS Chem. Biol. (2019)
Notes: As with all fluorescence microscopy, the photophysical properties of the fluorophore labels are crucial for SMLM. These are summarized in Tables 1 and 2 and include the laser excitation wavelength (λex), the fluorescence emission maximum (λem), the extinction coefficient at λmax (ε), and the fluorescence quantum yield (Φ). Properties specific to fluorophores activated with light include: the fluorescence contrast upon photoactivation (Fon/Foff), the turn-on (t1/2on) halftime, the photobleaching halftime (t1/2PB), and the photon yield of individual fluorophores before bleaching (N). The maturation halftime (t1/2M) is an important consideration for fluorescent proteins. For caged dyes, a critical factor is the photoactivation quantum yield (ΦPA).

## Slide 29
Jradi et al. ACS Chem. Biol. (2019)

## Slide 30: Current Approaches in SMLM
They differ primarily in how ON/OFF switching is achieved:
(f)PALM – photoactivation
STORM – photoswitching of activator and reporter dye-pair
dSTORM – conventional fluorescent probes in the presence of thiols transfer dyes to a long-lived OFF state
(f)BALM – binding and fluorescence activation of specific dyes
DNA-PAINT/Exchange-PAINT – transient oligonucleotide hybridization
Betzig et al. Science (2006), Hess et al. Biophy. J. (2006), Rust et al. Nature Methods. (2006), Heilemann et al. Angew Chem Int Ed Engl. (2008), Schoen et al. Nano Lett. (2011), Szczurek et al. Nucleic Acids Res. (2017), Jungmann et al. Nano Lett. (2010), Jungmann et al. Nat. Methods (2010), Jungmann et al. Nat. Methods (2014)

## Slide 31: Linkage errors of various fluorescent labelling approaches
Lelek et al. Nature Reviews Methods Primer (2021)
!!!
Notes: immunolabelling with primary (1st) and secondary (2nd) antibodies (yellow); labelling with a small camelid antibody (nanobody, light green), often in combination with green fluorescent protein (dark green), as shown, or when available directly binding to the protein of interest (not shown); labelling with a protein tag or genetically encoded protein, such as Eos (orange); direct labelling with a dye-conjugated ligand (such as the microtubule-binding compound docetaxel in this example); and incorporation of unnatural amino acids such as TCO*-lysine through genetic code expansion, which enables rapid labelling using functionalized synthetic dyes.

## Slide 32: Buffers in SMLM
Vogelsang et al. Angew Chem Int Ed Engl. (2008)
Different dyes blink optimally in different imaging buffers. Therefore, the acquisition of good quality 2- and 3-color images can be challenging.
Average photon counts per molecule in different buffers
Notes: Figure: Average photon counts per molecule a) for the red fluorophores Alexa647, MR121, ATTO647N, and Cy5, and b) for the green/yellow fluorophores Cy3B, ATTO565, and Alexa532 were obtained from at least three photobleaching movies using objective‐type total internal reflection microscopy. Measurements were carried out in PBS (black bars), oxygen‐depleted PBS with 1 mM AA (red), with 1 mM MV (blue), with ROXS (1 mM AA plus 1 mM MV (green), or 2 mM TX plus 1 mM MV (cyan), or 2 mM TX (magenta)).

## Slide 33: Buffers in SMLM
Advice: Once you have picked a dye for your SMLM experiment depending on your microscope (available lasers) and biological constraints, make yourself familiar with which buffers are most suitable and test & optimize them.

## Slide 34: Buffers in SMLM
Nahidiazar et al. PLOS ONE (2016)
Notes: RF EC24 endothelial cells were labeled for vimentin using a mouse monoclonal antibody (Ab) and secondary Ab labeled with Alexa dyes.

## Slide 35: Buffers in SMLM
Nahidiazar et al. PLOS ONE (2016)
Comparison of image quality in ageing Gloxy buffer (B) to that in OxEA buffer (A).
Advice: Always prepare your imaging buffer right before experiments &
adjust your imaging buffer choice for multicolor experiments.

## Slide 36: Overview
Basic principles 2D SMLM
Hardware for SMLM
Practical considerations: sample preparation, suitable dyes, linkage errors, and buffers
Processing, quantification, and interpretation of SMLM data
3D SMLM
Summary
Extra: New directions in SMLM
Extra: SRM as a multidimensional challenge

## Slide 37: Processing, quantification, and interpretation of SMLM data

## Slide 38: How large is an SMLM dataset?
50,000 frames taken with a 16-bit camera (1024 x 1024 pixel) lead to…
Total number of pixel on the detector:
1024 x 1024 = 1,048,576 pixel
Conversion into bytes:
16,777,216 bits / 8 = 2,097,152 bytes ~ 2 MB (one frame)
Total bits:
1,048,576 x 16 bits = 16,777,216 bits
Upscaled to 50,000 frames:
2 MB x 50,000 ~ 100 GB
That is a lot of data that needs to be handled…

## Slide 39: From raw SMLM data to reconstructed super-resolved image
Image Acquisition
Image at t1
Image at t2
Image at t3
t1 < t2 < t3
Betzig et al. Science (2006), Hess et al. Biophy. J. (2006), Rust et al. Nat Methods (2006)
Detection and localization
Render SMLM image
Widefield image
Superresolution image

## Slide 40: Detection
Approximate molecule location = Detection
Detection

## Slide 41: Detection and localization
Approximate molecule location = Detection
Accurate determination of the molecule center = Localization
Localization
Thompson et al. Biophysical Journal (2002)
vs.
Abbe Equation

## Slide 42: Error in localization – Equations
Thompson et al. Biophysical Journal (2002)
The modified equation to account for photon-counting noise, pixelation noise and background noise:
Error in localization
s Standard deviation of the PSF
N Number of photons collected
b Background noise
a Size of the pixel

## Slide 43: Detection and localization
Approximate molecule location = Detection
Accurate determination of the molecule center = Localization
Result: List of localizations (x,y coordinates, localization errors / uncertainties & add. parameters)
New images with localizations?
No, but a list of localizations!

## Slide 44: Pixel data vs. point clouds
In contrast to standard microscopy imaging that produces 2D pixel or 3D voxel grid data, SMLM generates big data of 2D or 3D point clouds with millions of localizations and associated uncertainties.
Therefore, processing, quantification, analysis and interpretation of SMLM must be approaches differently.
To this day, the SMLM data quantification and interpretation methods have yet to keep pace with the rapid advancement of SMLM imaging.
Khater et al. Biophysical Journal (2002)

## Slide 45: List of localizations
New images with localizations?
x,y – central coordinate
sigma - Gaussian width
Gaussian amplitude / offset
χ2 – goodness-of-fit
Uncertainty

## Slide 46: Sequence of SMLM processing and analysis steps
(4) Rendering
Super-resolved image (typ. 1028x1028)
(5) Quantification
Size (nm), Clusters, Stoichiometry
(3) De-drifting and Grouping
List of localizations (typ. 80,000)
(2) Sieving / Filtering
List of localizations (typ. 100,000)
(1) Detection & localization
Raw camera image (typ. 128x128x10,000)
List of localizations (typ. 200,000)

## Slide 47: High-density data / multiple emitters
B
A
A
B
Alternatively, you can filter/sieve your data to handle invalid localizations.
Advice: Choose suitable filtering parameters with the histogram of the respective fitting parameter.
Choosing a suitable multi-emitter fitting method/algorithm can help to deal with high-density data.

## Slide 48: SMLM’s resolution
So, when you filter/sieve out high uncertainties, you can reach insanely high resolutions, right?
You can filter high uncertainties, but at a certain point you’ll lose the structure’s integrity. The measure of resolution would be meaningless.
No, not really!
10.7
4.0
11.7
13.3
15.4
2.8
6.5
20.7
25.5
17.4
17.4
12.2
7.5
10.3
5.3
15.8
12.1
8.5
4.5
5.5
5.5
11.1
Uncertainty < 25 nm
4.0
2.8
6.5
7.5
5.3
8.5
4.5
5.5
5.5
Uncertainty < 10nm

## Slide 49: Drift-correction
0.5 µm
4 min movie
Drift becomes a limiting factor in achieving high resolution. Fortunately, it can be corrected for by fiducial-based or image-based (cross-correlation) strategies.
uncorrected
corrected

## Slide 50: Grouping / Blink-correction
Grouping radius rg
According to time
Frames
Intensity [AU]
According to space
Individual molecules
Grouped photon-averaged position
Grouping time tg
“Grouped” molecules
Notes: All localizations within a given region of space limited by grouping radius r_g and time limited by a grouping time t_g are considered to originate from the same molecule and their positions are averaged into a new grouped photon-weighted average molecular localization.

## Slide 51: Software for SMLM
Sage et al. Nature Methods (2019), Sage et al. Nature Methods (2015)

## Slide 52: Sequence of SMLM processing and analysis steps
(1) Detection & localization
(2) Sieving / Filtering
(3) De-drifting and Grouping
(4) Rendering
(5) Quantification
Raw camera image (typ. 128x128x10.000)
List of localizations (typ. 200.000)
List of localizations (typ. 100.000)
List of localizations (typ. 80.000)
Super-resolved image (typ. 1028x1028)
Size (nm), Clusters, Stoichiometry
With ThunderSTORM you can…
Ovesny et al. Bioinformatics (2014)

## Slide 53: Feel free to test it…
Software: Fiji plugin ThunderSTORM Plugin can be downloaded here: https://github.com/zitmen/thunderstorm/releases/tag/v1.3 and will be located here: Plugins > ThunderSTORM Online manual: https://github.com/zitmen/thunderstorm/wiki/Tutorials
Dataset: Tubulin 2D Long Sequence http://bigwww.epfl.ch/smlm/datasets/index.html?p=../challenge2013/datasets/Real_Long_Sequence
Ovesny et al. Bioinformatics (2014)
Use the Fiji plugin ThunderSTORM to analysis the Tubulin 2D Long Sequence Dataset

## Slide 54: Timeline of SMLM developments
Nicovich et al. Nature Protocols (2017)

## Slide 55: Spatial analysis approaches
Nicovich et al. Nature Protocols (2017)

## Slide 56: Fiji plugins for SMLM post-processing
Van der Linde et al. Journal of Physics D: Applied Physics (2019)

## Slide 57: Overview
Basic principles 2D SMLM
Hardware for SMLM
Practical considerations: sample preparation, suitable dyes, linkage errors, and buffers
Processing, quantification, and interpretation of SMLM data
3D SMLM
Summary
Extra: New directions in SMLM
Extra: SRM as a multidimensional challenge

## Slide 58: 3D SMLM

## Slide 59: Can we localize SM in the axial direction?
A microscope’s PSF is symmetric in the axial direction, which complicates single molecules' localization along the axial axes.
Simulated PSF
-200 nm
0 nm
+200 nm

## Slide 60: Decoding the axial information
Breaking the symmetry of the PSF enables us to decode axial information.
Astigmatism
Biplane
Double-helix
Cylindrical lens in detection path
Emitted light detected onto two cameras (or two halves of one camera) defocused relative to one another
Spatial light modulator / phase mask in detection path
Huang et al. Science (2008), Juette et al. Nature Methods (2008), Pavani et al. Optics Express (2008)

## Slide 61: Decoding the axial information
Lelek et al. Nature Reviews Methods Primer (2021)

## Slide 62: Decoding the axial information
Breaking the symmetry of the PSF enables us to decode axial information.
Shtengel et al. PNAS (2009)
iPALM
Resolution: 10-20 nm in all three dimensions

## Slide 63: Overview
Basic principles 2D SMLM
Hardware for SMLM
Practical considerations: sample preparation, suitable dyes, linkage errors, and buffers
Processing, quantification, and interpretation of SMLM data
3D SMLM
Summary
Extra: New directions in SMLM
Extra: SRM as a multidimensional challenge

## Slide 64: Summary

## Slide 65: Let’s wrap it up…
Merits
Very high resolution
Single molecule detection
Relative simple microscope setup
Can be combined with TIRF and inclined illumination (HILO)
Quantification of protein numbers
Upgrade solution for exciting setups to enable extended 3D localization using PSF engineering (e.g., Double Helix)
Image: Xu et al. Nature Methods (2012)

## Slide 66: Let’s wrap it up…
Special buffers/probes required
Not for thick samples (< 10 um)
Slow acquisition
Limited 3D (no sectioning)
Advanced post-processing needed
Prone to reconstruction artifacts
Structural resolution labeling density-dependent
Disadvantages
Image: Xu et al. Science (2013)

## Slide 67: Overview
Basic principles 2D SMLM
Hardware for SMLM
Practical considerations: sample preparation, suitable dyes, linkage errors, and buffers
Processing, quantification, and interpretation of SMLM data
3D SMLM
Summary
Extra: New directions in SMLM
Extra: SRM as a multidimensional challenge

## Slide 68: Thank you! Questions?
Ulrike Boehm, Ph.D.
ulrike.boehm@zeiss.com
Twitter: @ulrike_boehm
Opt to Work at ZEISS Job openings and applications
https://www.zeiss.com/corporate/int/careers/job-openings-and-applications.html
Photo: Martin Duckek

## Slide 69: New dimensions in SMLM

## Slide 70: MINFLUX emitter localization concept
Lelek et al. Nature Reviews Methods Primer (2021)
Notes: a | MINFLUX excitation concept for precisely probing emitter positions using minimal photon fluxes. A doughnut-shaped excitation beam (green) is moved sequentially to four probing positions r0, r1, r2 and r3 (coloured circles; probing range L) in the vicinity of a single fluorophore (orange star). If the doughnut centre coincides perfectly with the fluorophore position, no photons are emitted. The position of the fluorophore can be calculated with very high precision from the fluorescence photon counts (shown below). b | Example nuclear pore complexes imaged by MINFLUX.

## Slide 71: Expansion SMLM
Lelek et al. Nature Reviews Methods Primer (2021)
Notes: c | In expansion microscopy, samples are embedded in a gel that expands upon hydration. Immunolabelling of epitopes can be performed before or after gelation and expansion using linkers that bind to the gel and to a fluorescent dye. Full or partial protein digestion is commonly used to enable isotropic expansion. In order to enable direct stochastic optical reconstruction microscopy (dSTORM) imaging in photoswitching buffer, the sample is re-embedded in an uncharged polyacrylamide gel after expansion. d | Left: 3D post-labelling expansion dSTORM image of a 3.2× expanded and re-embedded sample showing 9-fold symmetry of the procentriole. Scale bar: 500 nm. Right: 3.1-fold expanded and re-embedded tubulin filaments, with magnified view of highlighted region. An xz side-view cross-section of a tubulin filament (bottom right) shows its hollow structure. Scale bars: 500 nm (vertical rectangle), 200 nm (small square).

## Slide 72: Deep learning-accelerated SMLM
Lelek et al. Nature Reviews Methods Primer (2021)

## Slide 73: SRM as a multidimensional challenge

## Slide 74: Inherent trade-offs in SRM
Schermelleh et al. Nature Cell Biology (2019)

## Slide 75: Decision tree for selecting SRM techniques
Schermelleh et al. Nature Cell Biology (2019)

## Slide 76: Appendix

## Slide 77: Switching buffers
Turkowyd et al. Anal. Bioanal. Chem. (2016)

## Slide 78: Switching buffers
Turkowyd et al. Anal. Bioanal. Chem. (2016)

## Slide 79: Multicolor SMLM
Turkowyd et al. Anal. Bioanal. Chem. (2016)

## Slide 80
Turkowyd et al. Anal. Bioanal. Chem. (2016)